For in vitro diagnostic use. Rx only.                                  This test has not been reviewed by the FDA

# INTENDED USE

The Quickkview  PLUS COVID-19 Antigen Test is a is a rapid  chromatographic immunoassay intended for the qualitative detection of nucleocapsid protein antigen from SARS-CoV-2 in nasal pharyngeal swab specimens collected from individuals who are suspected of COVID-19 by their healthcare provider within the first seven days of symptom onset.

Testing is limited to laboratories certified under the Clinical Laboratory

Improvement Amendments of 1988 (CLIA), 42 U.S.C. §263a, that meet the requirements to perform moderate, high or waived complexity tests. This test is authorized for use at the Point of Care (POC), i.e., in patient care settings operating under a CLIA Certificate of Waiver, Certificate of Compliance, or Certificate of Accreditation.

The Quickview PLUS COVID-19 Antigen Test does not differentiate between SARS-CoV and SARS-CoV-2.

Results are for the identification of SARS-CoV-2 nucleocapsid protein antigen. Antigen is generally detectable in upper respiratory specimens during the acute phase of infection. Positive results indicate the presence of viral antigens, but clinical correlation with patient history and other diagnostic information is necessary to determine infection status. Positive results do not rule out bacterial infection or co-infection with other viruses. The agent detected may not be the definite cause of disease.

Negative results, from patients with symptom onset beyond five days, should be treated as presumptive and confirmation with a molecular assay, if necessary, for patient management, may be performed. Negative results do not rule out COVID-19 and should not be used as the sole basis for treatment or patient management decisions, including infection control decisions. Negative results should be considered in the context of a patient’s recent exposures, history and the presence of clinical signs and symptoms consistent with COVID-19.

The Quickview PLUS COVID-19 Antigen Test is intended for use by medical professionals or trained operators who are proficient in performing rapid lateral flow tests.

# INTRODUCTION

Corona viruses are a large family of viruses that are common in many different species of animals, including camels, cattle, cats, and bats. The two highly pathogenic viruses, SARS-CoV and MERS-CoV, cause severe respiratory syndrome in humans, and the other four corona viruses (HCoV-NL63, HCoV-229E, HCoV-OC43 and HKU1) induce only mild upper respiratory diseases in immunocompetent hosts, although some of them can cause severe infections in elderly individuals1. COVID-19 is the virus associated with SARS-CoV-2. Corona viruses cause respiratory and intestinal infections in animals and humans. The virus is transmitted mainly via respiratory droplets that people sneeze, cough, or exhale. The incubation period for COVID-19 is currently estimated at between 2 and 14 days. Common symptoms of COVID-19 infection include fever, dry cough and respiratory symptoms such as shortness of breath and breathing difficulties. More serious cases develop severe pneumonia, acute respiratory distress syndrome, sepsis and septic shock that can lead to the death of the patient. People with existing chronic conditions seem to be more vulnerable to severe illness.  This COVID-19 Antigen Test is an immunoassay for the detection of nucleocapsid protein antigen from SARS-CoV-2..                                                 This test is based on a patented  technology which greatly enhances the speed, accuracy, and specificity of the canoniucal lateral flow.

# PRINCIPLE OF TEST

This device2, 3 is a chromatographic immunoassay. A progressive compression structure is built into the device to accelerate the reactions.

This device detects SARS-CoV-2 nucleocapsid protein antigen through visual interpretation of color lines.

Each device has a nitrocellulose membrane strip coated with antibodies against COVID-19 proteins as the test line (T).   The strip is also coated with Goat x Rabbit IgG, acting as an internal control, the control line (C). Another critical component is the conjugate pad with Colloidal gold-labeled anti-COVID-19 antibodies. A bottle with lysing buffer is included in this kit to lyse and deactivate the virus. If the COVID-19 virus is present in the specimen, the lysed proteins of COVID-19 will bind to anti-COVID-19 antibodies, forming antibody-antigen complexes and then move across the membrane surface pushed via progressive compression force.  If COVID-19 virus is present in the specimen, the T line will bind to the antibody-antigen complexes, and

the T line will appear in burgundy color. If there is no COVID-19 virus in the specimen, no burgundy T line will appear. The C line should always be visible after testing as a confirmation that the test is working properly.

During testing, the progressive structure progressively forces the specimen to mix with conjugated compounds thoroughly, and then flow onto the surface of the membrane strip which mobilizes the colored antibody conjugates. The reaction time can be as quick as three to seven (3-7) minutes.

# REAGENTS AND MATERIALS SUPPLIED

- Individually pouched test devices (20 pcs)

- Swab sample extraction tubes (20 pcs): Containing liquid extraction reagent

- Sterile nasal swabs (20 pcs)

- Instruction for Use (1)

- Quick Reference Guide (1)

# MATERIALS REQUIRED BUT NOT PROVIDED

- Timer

- Quickview COVID-19 Antigen Control Set*

(Available as separate purchase)

# TEST STORAGE AND STABILITY

The intact test kit can be stored at 39°-86°F (4°-30C) until the expiration date printed on the labels.

Exposing the kit to temperatures over 86°F (30°C) may reduce the shelf life or cause malfunction of the device.

# SPECIMEN PRESERVATION

- Specimens should be tested as soon as possible after collection; or

- Stored at 2°-8°C for 72 hours; or

- Stored at -20°C for up to 3 months; or

- Stored at -80°C for at least 1 year;

- Avoid repeat freezing and thawing.

# PRECAUTIONS

- For in vitro diagnostic use.

- For prescription use only.

- This test has been authorized only for the detection of proteins from SARS-CoV-2, not for any other viruses or pathogens.

- The instructions must be followed exactly to obtain accurate results.

- Do not remove the test from its sealed pouch until prior to use.

- Do not use expired devices and reagents.

- Use appropriate precautions in the collection, handling, storage, and disposal of patient samples and used kit contents.

- Use of Nitrile, Latex (or equivalent) gloves is recommended when handling patient samples.

- Inadequate or inappropriate sample collection, storage, and transport may yield false test results.

- Sample collection and handling procedures require specific training and guidance.

- To obtain the most sensitive results, directly test patient specimens without transport media.

- To obtain accurate results, do not use visually bloody or overly viscous samples.

- Testing should be performed in an area with adequate ventilation.

- Dispose of containers and unused contents in accordance with Federal, State, and Local regulatory requirements.

- Wear suitable protective clothing, gloves, and eye/face protection when handling the contents of this kit.

- Wash hands thoroughly after handling.

For in vitro diagnostic use. Rx only.                                  This test has not been reviewed by the FDA

# TEST COMPONENTS

# SPECIMEN COLLECTION

# IMPORTANT: Proper specimen collection, storage, and transport are critical for the performance of this device.

# Use standard clinical methods to collect nasal pharyngeal specimens.

- Nasal Pharyngeal Specimen:

Tilt the patient’s head back 70 degrees. Have the patient close his/her eyes as this helps minimize discomfort. Gently insert the swab through one of the nostrils and horizontally into the nasal passage up to the measured distance on the swab shaft or until resistance is met. Rotate the swab 2-3 times and then hold the swab in place for 5-10 seconds for better specimen absorption.

# SPECIMEN TRANSFER TO EXTRACTION TUBE

# 1. Remove the white cap from tube.

TIP

TIP

CAP

# 2. After collecting specimens with swab, insert the swab head into the liquid in extraction tube and swirl a minimum of three (3) times then leave in place for one (1) minute.

TIP

# 3. When removing swab, squeeze the sides of extraction tube pressing the swab head against the inside of extraction tube to extract as much liquid as possible into the tube. Then close the dropper tip and tightly seal the extraction tube with specimen.

# TEST PROCEDURE

- Prior to testing, specimen and all components of the kit must be equilibrated to room temperature.

- Remove the cassette from the foil pouch.  Do not open pouch until ready to perform the test.

NOTE:  Do not pull the cap off the device.

CAP

- Squeeze the extraction tube and apply two (2) drops vertically into the sample well of the cassette.

- Wait until the Burgundy color flow appears in the Result Window near the T line before closing the cap.

- Close the cap of test device with force until it clicks. Start the timer. Read the result in three to seven (3-7) minutes.

CAP

CAP

For in vitro diagnostic use. Rx only.                                  This test has not been reviewed by the FDA

# INTERPRETATION OF RESULTS

IMPORTANT: The T line should always be interpreted independently from the C line. Do not compare color intensities between lines. A faint test line also indicates a positive result.

Positive result

The appearance of the burgundy lines at both T and C positions indicates a positive result.

Negative result

The appearance of a burgundy line only at the C position indicates a negative result.

## Invalid result

If no line appears at C position, regardless of the appearance of the burgundy line at the T position, the result is invalid.

If the test is invalid, a new test should be performed with a new sample and a new extraction tube and test device.

Invalid results may relate to insufficient volume of specimen and/or incorrect testing procedure.  Review the procedure and repeat the test using a new test device.

If the problem persists, stop testing and contact customer services at 1-877-204-5071, 8am-5pm PST.

# QUALITY CONTROL

## Built-in Control

The Quickview® PLUS COVID-19 Antigen Test has a built-in (internal) procedural control, the C line. For daily quality control, it is recommended to record the result for each test run. The appearance of a burgundy red C line for COVID-19 Antigen Test indicates that the test has been performed correctly, including that the proper volume of specimen has been absorbed and the specimen flow has occurred. If the C line does not appear within 3 minutes, the test result is considered invalid.

External Controls

It is recommended that the external positive and negative controls be tested as a good laboratory practice to confirm the test procedure and to verify proper test performance.

*External controls, not included in this kit, may be purchased separately and used to verify that all reagents and procedures are performing properly.

The failure to obtain a negative result with the Negative Control or a positive result with the Positive Control likely indicates that the test was not performed properly or that the test reagents were not functioning properly.

# LIMITATIONS OF PROCEDURE

- Use of viral transport media may result in decreased test sensitivity, and directly testing specimens is recommended.

- The contents of this kit are to be used for the qualitative detection of SARS antigens from nasopharyngeal swab.

- This test detects both viable (live) and non-viable, SARS-CoV, and SARS-CoV-2. Test performance depends on the amount of virus (antigen) in the sample and may or may not correlate with viral culture results performed on the same sample.

- A negative test result may occur if the level of antigen in a sample is below the detection limit of the test or if the sample was collected or transported improperly.

- Failure to follow the Test Procedure may adversely affect test performance and/or invalidate the test result.

- Test results must be evaluated in conjunction with other clinical data available to the physician.

- Positive test results do not rule out co-infections with other pathogens.

- Positive test results do not differentiate between SARS-CoV and SARS-CoV-2.

- Negative test results are not intended to rule in other non-SARS viral or bacterial infections.

- Negative results, from patients with symptom onset beyond five days, should be treated as presumptive and confirmation with a molecular assay, if necessary for patient management, may be performed.

- If the differentiation of specific SARS viruses and strains is needed, additional testing, in consultation with state or local public health departments, is required.

# PERFORMANCE CHARACTERISTICS

Clinical Agreement Study

80 clinical samples were collected and evaluated in triplicate in the United States.  The clinical findings compared to a COVID-19 molecular assay are shown as follows:

| Method |  | PCR |  | Total Results |
| --- | --- | --- | --- | --- |
| COVID-19 Antigen Test | Results | Positive | Negative |  |
|  | Positive | 23 | 0 | 23 |
|  | Negative | 2 | 55 | 57 |
| Total Results |  | 25 | 55 | 80 |

Relative Sensitivity: 92%

Relative Specificity: 100%

REFERENCES

# REFERENCES

- Coronavirus Disease 2019 (COVID-19) https://www.cdc.gov/coronavirus/2019-ncov/index.html

- US patent 9,377,457, Progressive compression driven flow cartridge for analyte detecting strip and method.

- US patent 9,702,872, Rapid diagnostic test device by driven flow technology.

Targeting Systems

1453 N Cuyamaca St

El Cajon, CA 92020